## Slide 1: HFT Alignment Procedures
S. Margetis, KSU
Introduction
HFT Configuration – System hierarchy
(thoughts on) Proposed procedures
Tools/Implementation (Jonathan)
Plans, Timeline and Issues
Summary
Alignment procedures review, BNL, Oct. 12, 2012

## Slide 2
General Flowchart of Survey/Alignment Tasks
Sensor/Ladder Survey
Save – CVS?
Fit, get structures
Save – Db
GEANT geometry
Apply*
RawData
Apply
Alignment
Correct
Done
N
Y
Update – Db
HITS
Final Reco Pass
ForEachRun
ForEachNewSector/Ladder
Mis-align/Check Alignment software
* VMC, STV ready

## Slide 3: Introduction
Anything we build or touch or use needs Modeling, Survey and in-situ Alignment
  i.e. versioning
Survey will freeze position of sensors on sectors (PXL). Help also with sector on hemisphere (PXL?). For SSD/IST will freeze position of sensors on ladder and ladder shape
For each yearly Run the in-situ position of major detector elements needs to be rechecked

## Slide 4: Reference system hierarchy
STAR Magnet=STAR system
TPC system
Fixed x,y,z
Fixed-1D-rot
Float
Star Magnet defines overall system (Field map)
TPC is the first important system for HFT (relative positioning), attached to Magnet

## Slide 5
IDS/MSC
ESC/WSC attached to TPC wheel. It defines the HFT system’s relation (as a whole) to TPC system
See next slides for systems inside the HFT complex

## Slide 6
6
PST (IST)
OSC (SSD)

## Slide 7: General Layout
IST
IST
SSD

## Slide 8: Reference systems - comments
In general survey accuracy of critical components (relative pixel and/or sensor positions) expected to be better than minimum acceptable values
Will soon need surveyed positions of IDS targets
  to build ‘ideal’ position Db
  Sub-millimeter accuracies acceptable -> Tracks will fix them
All this information is represented as matrices (position/orientation) of their center-of-gravity. These matrices are used to define Local to Global transforms
GEANT geometry can/should be synchronized with Realistic Volume hierarchy instead of the current ‘patch-the-hit’ scheme
  - VMC environment will facilitate this

## Slide 9: Offline use of Geometry Info
y (up)
z (East)
x (South)
n,b
t,g
d,a
Definitions
Local-to-Global transforms are done in terms of TGeoHMatrix
General form is:
This can be e.g. the center of a sensor or a pixel.
n,d,t are unit vectors and b,a,g the corresponding rotation angles, RHS
TGeoHMatrix definition
Transform example

## Slide 10
WG = Tpc2Global * GL * SG * LS * WLL;
WaferInGlobal=Tpc2Magnet *SsdinTpc*SectorInSSD*LadderInSector*WaferInLadder
Local <-> Global transforms
WG = Tpc2Global * GL * LO * WLL;
WaferInGlobal=Tpc2Magnet *IDS2Tpc*Ladder2IDS*WaferInLadder
WG = Tpc2Global * GL * PI *LO * WLL;
WaferInGlobal=Tpc2Magnet *IDS2Tpc*PST2IDS*Ladder2PST*WaferInLadder
PG = Tpc2Global *GL * PI *DP * SD * WLL;
PXLInGlobal=Tpc2Magnet*IDS2Tpc*PXL2IDS*DShell2PXL*Sector2DShell*(Pxl-Sector)
OLD SSD/SVT
HFT SSD
HFT IST
HFT PXL

## Slide 11: Alignment methods (outline only)
There are ‘Global’ and ‘Self’ Alignment methods
  Global uses mostly ‘external’ track information
  Self uses mostly ‘internal’ track information
  For HFT will use a mix (more Self !)
We have successful ‘Global’ methods already in place (SVT/SSD)
  TPC distortions, t0, ‘track tof’ etc is a problem
We now have significant sensor overlap to make use of ‘Self’ alignment methods. We also have high precision PXL info with excellent sector rigidity, survey info, placement.
We need to use this advantage
We lack a hardware monitoring system. Once detectors are installed we rely on survey and alignment software

## Slide 12
TPC only tracks
  Global alignment of SSD (+SVT) with respect to TPC
  (Local) Alignment of SSD ladders:
TPC + SSD tracks
  (Global) Alignment of SVT Clam Shells
  (Local) Alignment of SVT ladders
TPC + SSD + SVT tracks
  Check consistency and
  re-evaluate SVT & SSD hit errors
Statistics needed:
1 mm  ~20 micron: reduction factor 50
~2,500 tracks per SVT sensor
data sample with ~250,000 tracks -> 250K CuCu events
SSD/SVT procedure (old)

## Slide 13
The sequence to be followed for each detector is:
1) SSD Alignment: (TPC tracks Only)
Global - SSD on Global and Sectors on Global;
Local - SSD Ladders on Sectors;
2) SVT Alignment: (TPC+SSD hits on tracks)
Global - SVT on Global and Shells on Global;
Local - SVT Ladders on Shells; (Drift Velocities);
3) Consistency Check: (TPC+SSD+SVT hits on tracks)
Global;Local (ladders);Drift Velocities;

## Slide 14: Procedure:
Remember: PXL detector is a big asset (avoid TPC distortions):
Global Alignment of PXL sectors
  Relative alignment of PXL halves using overlap region AND using Event vertex found by each half (hit level)
  Global positioning in STAR (TPC) using primary TPC tracks
Primary tracks with TPC+PXL hits
  Alignment of IST ladders with respect to PXL
Primary tracks with (All – SSD) hits
  Alignment of SSD ladders
Check
HFT
We assume that ladders on sectors and sectors in halves are pre-surveyed to specs
Other methods (Millepede) should be developed in parallel

## Slide 15: Figure of merit for HFT alignment.
For alignment we use “good” (well defined) tracks fitted with the primary vertex.
  Use of primary tracks significantly improves precision of track predictions in Silicon detectors and reduces influence of systematics.
In order to minimize TPC space-charge distortions (and PXL pileup) we will need to use low luminosity data
Pointing accuracy, aka Impact parameter resolution:
  DCA resolution (in bending XY  plane: DCA), and resolution in non-bending plane: z, is figure of merit
  2DCA= 2vertex+ 2track + 2MCS (the same for non-bending plane),
  primary vertex resolution: vertex ~ 3mm+(120 mm / √Nch); for central Au+Au collisions turns out to be ~5 mm
  track pointing resolution: track ~ 1.5 XY in our case, where XY is intrinsic detector precision (~10mm)  alignment errors,
  Multiple Coulomb Scattering (MCS): MCS ~ 20mm / bp (GeV/c) (for thin PXL)
  Overall mis-alignments of < 10 mm are acceptable (no big impact)
Notes: Heavy d

## Slide 16: Some results: DCA Resolution
The main factors determining the DCA resolution of the SVT/SSD is the mass (scattering) and the distance of the first layer from the vertex
The following figures show that we are close to the limits of the device which indirectly shows that Alignment/Calibration errors are a subset of the overall errors

## Slide 17: DCA resolution
Number of Silicon Points fitted to track | XY @1GeV/c (m) | Z @1GeV/c (m)
0 -  TPC only | 3350 | 1840
1 -  TPC+SSD | 967 | 993
2 -  TPC+SSD+SVT | 383 | 351
3 -  TPC+SSD+SVT | 296 | 232
4 -  TPC+SSD+SVT | 281 | 212
DCA resolution = standard deviation of global track DCA with respect to the primary vertex.
With increasing no. of fitted Si points it is improved by ~ order of magnitude.
Contribution from tracking (constant term) is comparable with MCS @ 1 GeV/c  reached the goal !
XY
Z

## Slide 18: SVT and SSD resolutions after Calibration&Alignment
Calibration&Alignment procedure for Run V has been done for 3 data samples : 62 GeV FF, 200 GeV RF and FF
Resolution has been estimated from hit pull analysis.
Results (average over 3 samples) are:
  SVT:
     = 49±5 m
    Z = 30±7 m
  SSD resolution:
     = 30 m (set to design value since << MCS)
    Z = 742 ± 41 m

## Slide 19: Procedure
TPC only tracks
  Global alignment of SSD (+SVT) with respect to TPC
  (Local) Alignment of SSD ladders:
TPC + SSD tracks
  (Global) Alignment of SVT Clam Shells
  (Local) Alignment of SVT ladders
TPC + SSD + SVT tracks
  Check consistency and
  re-evaluate SVT & SSD hit errors
Statistics needed:
1 mm  ~20 micron: reduction factor 50
~2,500 tracks per SVT sensor
data sample with ~250,000 tracks -> 250K CuCu events
PXL detector is a big asset (avoid TPC distortions):
Primary tracks with TPC+PXL hits
  Relative alignment of PXL halves (check survey)
  Alignment of IST ladders with respect to PXL
Primary tracks with (All – SSD) hits
  Alignment of SSD ladders
OLD ->
NEW ->

## Slide 20: Open Items/Issues
Need to finalize the PXL sensor representation in Db. This will be done with the survey of the prototype sector
Need to setup Data formats, Db(s), code to deliver matrices
Need to clarify/verify/define reference system (physical) dependencies and hierarchy
Need to know/map the (realistic) error of every survey step
Need to start simulations to determine alignment software performance
Need to rework GEANT geometry synchronization (STV, VMC)
Need to finalize SSD procedures and initialize/define IST ones
Need to include gravitational sagging in SSD and IST (?) model
Need to keep/use expertise around
Need to rework/prioritize Software Summer activities
See also sub-system specific issues

## Slide 21: SURVEY
Survey was performed for both SSD and SVT
For the SVT we got information about:
  Wafers on Ladder (High precision [<1micron] Nikon camera)
  Ladders on Clamshells (~25 micron accuracy)
No survey info for relative Clamshell placement
No survey in situ
No re-survey after water leakage or ladder replacement
No hardware position monitoring in situ
No cosmics or Z0
Only wafer position on ladder used. The rest just as a starting point for software

## Slide 22: Procedure (further details)
  SVT drift velocity: the first approximation of SVT drift velocity is obtained from t_min, t_max fits for each hybrid.
  TPC only tracks
    Global alignment of SSD (+SVT) with respect to TPC
    (Local) Alignment of SSD ladders: ladders translations up to ~200 m and rotations (especially around y-axis) of up to ~20mrad. After fine tuning the majority had translations of < 20 m and rotations <0.5mrad, all within errors.
  TPC + SSD tracks
    (Global) Alignment of SVT Clam Shells
    (Local) Alignment of SVT ladders
    Correction to SVT drift velocities. SVT drift velocities have been refitted including extra dependence on drift distance and anode (up to 3rd degree Tchebyshev). This fit reduced hit residuals from ~100 m to ~10 m.
  TPC + SSD + SVT tracks
    Check consistency and
    re-evaluate SVT & SSD hit errors
Statistics needed:
  1 mm  ~20 micron: reduction factor 50
  ~2,500 tracks per SVT sensor
  data sample with ~250,000 tracks -> 250K CuCu events

## Slide 23: Alignment procedure
Use:
  Notations : Global X, Y,Z, Local u ≡ x (drift), v ≡ z, w ≡ y, α,β,γ are rotation around u, v, w or global X,Y,Z, respectively. Units are cm.
  Rigid body model has been applied (ignore possible twists effects, etc for the moment).
  A misalignment model (D0 alignment model): Taylor’s expansion with respect to misalignment parameters (3D shifts (Δu,Δv,Δw) and 3D rotations(Δα, Δβ, Δγ)) for deviations of measured hit position from predicted (from other detectors) primary track position on a measurement plane
  A misalignment parameter has been calculated as a slope with straight line fit of histogram of most probable values for above deviations versus corresponding track coordinates or inclination to detector plane (see examples below)
Calibration sample:
  ~250 Kevents of Cu+Cu at 62 GeV/nucleon at the end of fills (low luminosity, low space charge → low TPC distortions). TPC is drifting too!
mkm=micron

## Slide 24: Methods (alignment) II
For alignment we use “good” (with well defined parameters) tracks fitted with the primary vertex.
  Use of primary tracks significantly improves precision of track predictions in Silicon detectors and reduces influence of systematics.
Precision of the method is checked with simulation (blind)
  Accuracy ~10 m in detector position and ~0.1 mrad in its rotation.
There is a problem when we start far from minimum because there are significant correlations among alignment parameters.
To solve this problem as a starting point we use Least-Squares Fit with above derivatives to get first approximation for the parameters.
  The precision of this method is less than slopes method but it does provide a reasonable approximation to use slopes.
Notes: Heavy d

## Slide 25: Step 1) Global SVT, SSD Alignment using Survey and TPC info
The SVT Clam-Shells and SSD Sectors were aligned using TPC tracking info only.
  SVT and SSD ladder survey info was used/assumed at this point. Ladder-on-Shell accuracy of survey data estimated (hard/soft) 20-30 mkm
  Proper math for Global shifts/rotations was developed (same procedure as in local)
  Procedure checked for accuracy and limitations with Monte Carlo blind tests.
In order to avoid drift velocity effects in the SVT, only the first 4mm of drift were used (|u| in range [2.5,2.9] cm out of [0.0,3.0] cm total Si drift)
Also excluded 1mm around readout anodes (due to variations in the focusing electric fields surrounding the anodes)
When done Shells/Sectors were (on average) aligned to better than ~50 mkm in translations, and a few mrad in rotations.
PTO ->

## Slide 26
An example of results can be seen below and at: http://www.star.bnl.gov/STAR/comp/reco/SVT/Alignment/Pass37_TpcOnly/C/Global/
and http://www.star.bnl.gov/STAR/comp/reco/SVT/Alignment/Pass37_TpcOnly/C/Results.Sector_5FriApr2818:40:172006Pass37_TpcOnly_CPlotsG44GNowafersNsp_u_2.5-2.9NFP25rCut0.5cm
β is rotation around Y axis

## Slide 27: Step 2) SSD Ladder tuning using TPC info
Although SSD Sectors were good on the average, individual Ladders showed translations up to ~200mkms and rotations (especially around y-axis) of up to ~20mrad. A fine-tuning was performed.
http://www.star.bnl.gov/STAR/comp/reco/SVT/Alignment/Pass37_TpcOnly/C/Results.Pass37_TpcOnly_CPlotsG44GNoWafers_u_2.5-2.9NFP25rCut0.5cmFriApr2819:39:262006
After the SSD Ladder fine tuning the majority had translations of <20mkm and rotations <0.5mrad, all within errors.
http://www.star.bnl.gov/STAR/comp/reco/SVT/Alignment/Pass37_TpcOnly/J/Ssd/Results.Pass37_TpcOnly_JPlotsG44G_u_2.5-2.9NFP25rCut0.5cmSunApr3022:38:302006
(see also next slide)
After this step the SSD Geometry was frozen and SSD info was put on tracking with proper errors (specs) (200 mkm or R-Phi and 700 mkm in Z)

## Slide 28
BEFORE
AFTER
Example of correcting a SSD individual ladder rotation around the z-axis

## Slide 29: Step 3) SVT Ladder Z-tuning using TPC+SSD info
Although SVT Shells, as a whole, were good on the average, individual Ladders showed Z-translations up to ~400mkms (but the bulk around 100mkms). We believe that this discrepancy between survey and in-situ positions is due to work done on Shells after the survey was completed (water pipe leakage). Also 2 Ladders were replaced and serviced.
Touching the detector after the survey is done should be avoided
After the SVT Ladder fine Z-tuning the majority has translations of <20mkm
http://www.star.bnl.gov/STAR/comp/reco/SVT/Alignment/Pass49_Q/Ladders
See next slide for example

## Slide 30
BEFORE
AFTER
Example of fine tuning the z position of an SVT ladder using TPC+SSD info

## Slide 31: HFT Internal (Self) Alignment
  An iterative method on track/vertex fitting
    The SVT/SSD hits are associated with tracks using the TPC tracks and then fitted.
    The event vertex is determined, the tracks refitted with the vertex and the hit residuals determined
    A correction is determined and the process starts again with the new hit positions
    Initial results encouraging
  The Millipede code was also tried as is
    Problem of strong correlation of parameters is still not resolved
    A modified version of this approach is currently under investigation

## Slide 32: Summary
Coming

## Slide 33: References
“The STAR time projection chamber a unique tool for studying high multiplicity events at RHIC”, M.Anderson et al., NIM A499: 652,2003.
“The laser system for the STAR time projection chamber”, J. Abele et al., NIM A499: 692,2003.
“Correcting for distortions due to ionization in the STAR TPC”, G. Van Buren et al.,NIM A566:22-25,2006.
“The STAR Silicon Vertex Tracker” A large area Silicon Drift Detector”, R.Bellwied et Al., NIM A499: 640, 2003.
“The STAR silicon strip-detector (SSD)”, L.Arnold et al., NIM 2003 A499: 652, 2003.
“Alignment Strategy for the SMT Barrel Detectors”, D.Chakborty, J.D.Hobbs, October 13, 1999. D0 Note (unpublished)
“Sensor Alignment by Tracks”, V.Karimaki et al.,CMS CR-2004/009 (presented at CHEP 2003)
http://phys.kent.edu/~margetis/STAR/HFT/Survey/SVTSmallScaleSelfAlignment.pdf
http://phys.kent.edu/~margetis/STAR/HFT/Survey/SVT_Alignment_JPCSL.pdf

## Slide 34
http://drupal.star.bnl.gov/STAR/comp/calib/svt/selfalign

## Slide 35
D.Chakraborty, J.D.Hobbs, D0 note Oct.13, 1999